## Laboratory Information
Field | Value
Laboratory Name | Embu PGH
Laboratory Number | 25985
Laboratory Address | 169-40100, Ahero
Laboratory Telephone | 87587999
Laboratory Fax | 02359875
Laboratory Email | aherolab@gmail.com
Head of Laboratory | Otieno Owang
Telephone (Head of Laboratory - Personal) | 071336985
Telephone (Head of Laboratory - Work) | 0722698598
Lab Level | June Alice_County
Lab Affiliation | GOK

## Staffing Summary
Profession | Employees | Adequate
Degree-holding Professional Staff | 5
Diploma-holding Professional Staff | 5
Certificate-holding Professional Staff | 2
Microscopist
Data Clerk | 1
Phlebotomist | 2
Cleaner | 1
Is the cleaner(s) dedicated to the laboratory only?
Has the cleaner(s) been trained in safe waste handling?
Driver | 1
Is the driver(s) dedicated to the laboratory only?
Has the driver(s) been trained in biosafety?
Other

## SLMTA Information
Field | Value
Official SLMTA Audit? | No
Audit Start Date | 2017-06-05
Audit End Date | 2017-06-05
Name(s) and Affiliation(s) of Auditor(s) | judith Malonza, kemunto agata, June Alice
Type of SLMTA Audit | Baseline Audit
Total # of tests before SLMTA | 15
Total # of tests this year | 28
Cohort ID | 0
SLMTA Labtype | June Alice- County
Date of baseline audit | 0000-00-00
Date of SLMTA workshop #1 | 2017-03-06
Date of exit audit performed | 0000-00-00
Baseline Score | 0
Stars obtained | Not Audited
Exit Score | 0
Stars obtained | Not Audited
Date of Last Audit | 0000-00-00
Score for Last Audit | 0
Prior Audit Status | Not Audited

## Assessment Details
Question | Description | Response | Notes
1 | 1.1 Legal Entity | Yes
2 | 1.2 Laboratory Quality ManualIs there a current laboratory quality manual, composed of the quality management system’s policies and procedures, and has the manual content been communicated to and understood and implemented by all staff?
3 | a) Quality policy statement that includes scope of service, standard of service, measurable objectives of the quality management system, and management commitment to compliance. | Yes
4 | b) Documented policies for the quality management system that meet the requirements of ISO15189:2012 <strong>(Refer to Question 1.5 of this checklist for list of policies required)</strong> | No | 232
5 | c) Description of the quality management system and the structure of its documentation | Yes
6 | d) Reference to supporting procedures (SOPs), including managerial and technical procedures | Yes
7 | e) Description of the roles and responsibilities of the laboratory director, or laboratory manager, quality manager, and other key personnel (laboratory to define its key personnel) responsible for ensuring compliance | Yes
8 | f) Records of review and approval of the quality manual by authorized personnel | No | to to to
9 | g) Records to show that the quality manual was communicated to and understood by the lab personnel | Yes
10 | 1.3 Document and Information Control System | Yes
11 | 1.4 Document and Records | Yes
12 | 1.5 Laboratory Policies and Standard Operating ProceduresAre policies and standard operating procedures (SOPs) for laboratory functions current, available and approved by authorized personnel?
13 | Ethical Conduct | Yes
14 | Document Control | Yes
15 | Control of Records | Yes
16 | Communication (internal and external) | No | due to this
17 | Service Agreements | Yes
18 | Examination by Referral Laboratories and Consultants | Yes
19 | External Services and Suppliers | Yes
20 | Purchasing and Inventory Control | Yes
21 | Advisory Services | Yes
22 | Resolution of Complaints and Feedback | Yes
23 | Identification and Control of Nonconformities | Yes
24 | Corrective Action | Yes
25 | Preventive Action | Yes
26 | Continual Improvement | Yes
27 | Control of Records | Yes
28 | Internal Audits | Yes
29 | Risk Management | Yes
30 | Management Review | Yes
31 | Personnel Management | Yes
32 | Personnel Training | Yes
33 | Competency Assessment | Yes
34 | Authorization | Yes
35 | Review of Staff Performance | Yes
36 | Accommodation and Environmental Conditions | Yes
37 | Laboratory Equipment | Yes
38 | Calibration of Equipment | Yes
39 | Pre-examination Process | Yes
40 | Validation and Verification of examination procedures / Equipment | Yes
41 | Measurement Uncertainty | Yes
42 | Biological Reference Intervals or Clinical Decision Values | Yes
43 | Documentation of examination procedures | Yes
44 | Laboratory Contingency Plan | Yes
45 | Quality Control Quality Assurance | Yes
46 | Reporting and Release of Results | Yes
47 | Laboratory Information System (LIS) (Computerized or non-computerized) | Yes
48 | Laboratory Safety Manual | Yes
49 | 1.6 Policy and SOPs Accessibility | Yes
50 | 1.7 Policies and SOPs Communication | Yes
51 | 1.8 Document Control Log | Yes
52 | 1.9 Discontinued Policies and SOPs | Yes
53 | 1.10 Data Files | Yes
54 | 1.11 Archived Results Accessibility | Yes
55 | 2.1 Routine Review of Quality and Technical RecordsDoes the laboratory routinely perform a documented review of all quality and technical records? Does the laboratory review include the following?
56 | a) Follow-up of action items from previous reviews | Yes
57 | b) Status of corrective actions taken and required preventive actions | No | No action items
58 | c) Reports from personnel | Yes
59 | d) Environmental monitoring log sheets | Partial | Not reviewed consistently
60 | e) Specimen rejection logbook | Yes
61 | f) Equipment calibration and maintenance records | Yes
62 | g) IQC records across all test areas | Yes
63 | h) Outcomes of PTs and other forms of Inter-laboratory comparisons | Yes
64 | i) Quality indicators | Yes
65 | j) Customer complaints and feedback | Yes
66 | k) Results of improvement projects | Yes
67 | l) Documentation of review and action planning with staff for resolution and follow-up review | No | Not documented
68 | 2.2 Management ReviewDoes the laboratory management perform a review of the quality system at a management review meeting at least annually?
69 | Review Input: Does the management review meeting include the following inputs? | Yes
70 | b) Assessment of user feedback | Yes
71 | c) Staff suggestions | Yes
72 | d) Internal audits | Yes
73 | e) Risk management | Yes
74 | f) Use of quality indicators | No | Not used
75 | g) Assessments by external organizations | Yes
76 | h) Results of participation in inter-laboratory comparison programmes (PT/EQA) | Yes
77 | i) Monitoring and resolution of complaints | Partial | not done consistently
78 | j) Performance of suppliers | Yes
79 | k) Identification and control of nonconformities | Yes
80 | l) Results of continual improvement including, current status of corrective actions and preventive actions | Yes
81 | m) Follow-up actions from previous management reviews | Yes
82 | n) Changes in the volume and scope of work, personnel, and premises that could affect the quality management system | Yes
83 | o) Recommendations for improvement, including technical requirements | Yes
84 | p) Review of quality objectives and the quality policy for appropriateness and continuous improvement | Yes
85 | Review Output: Does the management review meeting include the following outputs? | Partial | ddddhu
86 | b) Does the output records of the MR meeting capture decisions made, persons responsible for actions to be taken and timeframes? | Yes
87 | c) Does the report address resources required (human, financial, material)? | Yes
88 | d) Does it refer to improvement for the users? | Yes
89 | e) Does it refer to improvement of the effectiveness of the quality system? | Yes
90 | f) Were the quality objectives and the quality policy reviewed for appropriateness and continuous improvement? | Yes
91 | 2.3 Are findings and actions from MR communicated to the relevant staff? | Yes
92 | 2.4 Does lab management ensure actions from MR are completed within defined timeframes? | Yes
93 | 3.1 Duty Roster And Daily Routine | Yes
94 | 3.2 Organizational Chart and External/Internal Reporting Systems | Yes
95 | 3.3 Laboratory DirectorIs the laboratory directed by a person(s) with the competency, delegated responsibility to perform the following;
96 | a) Provide effective leadership, budgeting and planning | No | no budgets
97 | b) Communicate with stakeholders | Yes
98 | c) Ensure adequate competent staff | Yes
99 | d) Ensure the implementation of the QMS | Yes
100 | e) Selection and monitoring of lab supplies | Yes
101 | f) Selection and monitoring of referral labs | No | not done
102 | g) Ensure a safe lab environment | Yes
103 | h) Advisory services | Yes
104 | i) Provide professional development programs for laboratory staff | Yes
105 | j) Address complaints, requests or suggestions from staff and/or lab users | Yes
106 | k) Design and implement a contingency plan | Yes
107 | 3.4 Quality Management System OversightIs there a quality officer/manager with delegated responsibility to oversee compliance with the quality management system?
108 | a) Is there an appointment letter, job description available or terms of reference? | Yes
109 | b) Does the quality manager ensure that processes needed for the quality management system are established, implemented, and maintained? | Yes
110 | c) Does the QM report to management at which decisions relating to quality are made? | Yes
111 | d) Does the QM promote awareness of users’ needs and requirements throughout the organization? | Yes
112 | e) Does the QM participate in management reviews? | Yes
113 | 3.5 Personnel Filing SystemAre records of personnel maintained and do they include the following?
114 | a) Educational and professional qualifications | Yes
115 | b) Copy of certification or license to practice, when applicable | Yes
116 | c) Previous work experience e.g. CV | Yes
117 | d) Job descriptions | Yes
118 | e) Introduction of new staff to the laboratory environment | Yes
119 | f) Training in current job tasks including vendor training received on -site | Yes
120 | g) Competency assessments | Yes
121 | h) Records of continuing education | Yes
122 | i) Reviews of staff performance | Yes
123 | j) Reports of accidents and exposure to occupational hazards | Yes
124 | k) Immunization status, as applicable relevant to assigned duties | Yes
125 | l) Letter of employment or appointment | Yes
126 | m) Employee medical surveillance records | No | No records of medical surveillance
127 | 3.6 Laboratory Staff TrainingIs there a system for training that covers the following?
128 | a) The quality management system | Yes
129 | b) Assigned work processes, procedures and tasks | Yes
130 | c) The applicable laboratory information system | Yes
131 | d) Health and safety, including the prevention or containment of the effects of adverse incidents | Yes
132 | e) Laboratory Ethics | Yes
133 | f) Confidentiality of patient information | Yes
134 | g) Is there supervision for persons undergoing training | Yes
135 | h) Continuous medical education | Yes
136 | i) Review of effectiveness of the training program | Yes
137 | 3.7 Staff Competency Assessment and retrainingIs there a system for competency assessment that covers the following?
138 | a) Are competency assessments performed according defined criteria | Yes
139 | b) New hires | No | not documented
140 | c) Existing staff | Yes
141 | d) Retraining and re-assessment where needed | Yes
142 | 3.8 Staff MeetingsAre staff meetings held regularly and do the meetings address the following items?
143 | a) Follow-up of action items from previous staff meeting | Yes
144 | b) Systemic and or recurrent problems and issues addressed, including actions to prevent recurrence | Yes
145 | c) Complaints | Yes
146 | d) Communication on reviewed/revised/redundant SOPs | Yes
147 | e) Review of results from prior corrective actions | Yes
148 | f) Discussion and evaluation of improvement topics/projects | Yes
149 | g) Feedback given by staff that have attended hospital meetings, external meetings, training, conferences, workshops, etc. | No | no feedback
150 | h) Relay of reports and updates from lab attendance at meetings with clinicians (the use of lab services and/or attendance at clinical rounds) | Yes
151 | i) Recording and monitoring of meeting notes for progress on issues. | No | nor done
152 | 4.1 Advice and Training by Qualified Staff | Yes
153 | 4.2 Resolution of Complaints | Yes
154 | 4.3 Laboratory Handbook for Clients - information to users | Yes
155 | 4.4 Communication Policy on Delays in Service | Yes
156 | 4.5 Evaluation Tool and Follow up | No | No evaluation tool
157 | 5.1 Adherence to Proper Equipment Protocol | Yes
158 | 5.2 Equipment Operation | Yes
159 | 5.3 Equipment and Method Validation/Verification and DocumentationAre newly introduced equipment and methods validated/verified on-site and is documented evidence available?
160 | a) Are specific verification/validation protocols in place for each equipment and examination procedure? | Yes
161 | b) Is validation performed for all laboratory designed or developed methods, standard methods used outside their intended scope and validated methods that are subsequently modified? | Yes
162 | c) Has validation information been obtained from the manufacturer/method developer as part of the verification? | Yes
163 | d) Have performance characteristics been appropriately selected and evaluated as per intended use? | Yes
164 | e) Were the verification/validation studies appropriate and adequate? | Yes
165 | f) Was the analysis of data appropriate for the selected performance characteristics? | Yes
166 | g) Have the verification/validation results/reports been reviewed and approved by an authorised person? | Yes
167 | 5.4 Measurement uncertainty of measured quantity testsDoes the laboratory have documented estimates of measurement of uncertainty (UM)?
168 | a) Has the laboratory calculated the measurement uncertainty for each quantitative measurement procedure? | Yes
169 | b) Has the laboratory defined the performance requirements (factors that affect the UM) for the measurement uncertainty of each measurement procedure and regularly review estimates of measurement uncertainty? | No | Not documented
170 | c) Does the lab make its calculated measurement of uncertainty available to its users upon request? | No | Not documented
171 | 5.5 Equipment Record MaintenanceIs current equipment inventory data available on all equipment in the laboratory?
172 | a) Name of equipment | Yes
173 | b) Manufacturer’s or authorized supplier contact details | Yes
174 | c) Condition received (new, used, reconditioned) | Yes
175 | d) Serial number | Yes
176 | e) Date of receiving | Yes
177 | f) Where equipment is obsolete, date when put 'out of service' | Yes
178 | g) Date of entry into service(after validation / verification) | Yes
179 | h) Location | Yes
180 | 5.6 Equipment Maintenance RecordsIs relevant equipment service information readily available in the laboratory?
181 | a) Service contract information | Yes
182 | b) Contact details for service provider | Yes
183 | c) Decontamination Records | Yes
184 | d) Engineer or service provider preventative maintenance records | Yes
185 | e) Last date of service | Yes
186 | f) Next date of service | Yes
187 | 5.7 Defective Equipment Waiting for Repair | No | Broken down centrifuge not labeled
188 | 5.8 Obsolete Equipment Procedures | Yes
189 | 5.9 Equipment Calibration and Metrological traceability ProtocolIs routine calibration of laboratory equipment (including pipettes, centrifuges, balances, and thermometers) scheduled, as indicated on the equipment, and verified?
190 | a) Is routine calibration of laboratory ancillary equipment (including pipettes, centrifuges, balances, and thermometers) scheduled, at minimum following manufacturer recommendations and verified? | Yes
191 | b) Is the calibration traceable (e.g. use of reference materials and equipment like certified thermometers, tachometer? | Yes
192 | c) Is there evidence of review of calibrations certificates/results by the laboratory before acceptance back into use? | Yes
193 | d) Is certified reference materials, examination and calibration by another procedure, use of mutual consent standards or methods used for in house calibrations? | Yes
194 | 5.10 Equipment Preventive Maintenance | Yes
195 | 5.11 Equipment Service Maintenance | Yes
196 | 5.12 Equipment Malfunction - Response and Documentation | Yes
197 | 5.13 Equipment Repair Monitoring and DocumentationAre repair orders monitored to determine if the service is completed? Does the laboratory verify and document that it is in proper working order before being put it back into service?
198 | a) Are repair orders monitored to determine if the service is completed? | Yes
199 | b) Does the laboratory verify and document the equipment is in proper working order before being put it back into service? | Yes
200 | 5.14 Equipment Failure - Contingency Plan | Yes
201 | 5.15 Manufacturer's Operator Manual | No | Water bath manual missing
202 | 5.16 Laboratory Testing Services | Yes
203 | 6.1 Internal AuditsAre internal audits conducted at intervals as defined in the quality manual and do these audits address areas important to patient care?
204 | a) Is there an audit plan/schedule that ensures all activities of the QMS are audited? | Yes
205 | b) Are audits being carried with minimal conflict of interest e.g. where possible, carried out by persons who are not involved in lab activities in the section being audited? | Yes
206 | c) Are the personnel conducting the internal audits trained with proven competency in auditing managerial and/or technical requirements? | No | no certs
207 | d) Is cause analysis performed for nonconformities/noted deficiencies? | Yes
208 | e) Are internal audit findings documented and presented to the laboratory management and relevant staff for review? | Yes
209 | 6.2 Audit Recommendations and Action Plan & Follow up
210 | a) Are internal audits reports generated? | Yes
211 | b) Are recommendations for corrective/preventive actions made based on audit findings? | Yes
212 | c) Is an action plan developed with clear timelines, assigned personnel & documented follow-up within the timeframe defined by the laboratory? | Yes
213 | 6.3 Risk ManagementAre assessment of potential pitfalls performed for all laboratory processes including pre examination, examination and post examination?
214 | a) Documented assessment of potential pitfalls for all processes | Yes
215 | b) Documented actions taken to reduce or eliminate identified potential pitfalls | Yes
216 | 7.1 Inventory and Budgeting System | Yes
217 | 7.2 Specification for supplies/consumables | Partial | only one evaluated
218 | 7.3 Service Supplier Performance Review | Yes
219 | 7.4 Inventory ControlDoes the lab maintain records for each reagent and consumable that contributes to the performance of examinations. These records shall include but not be limited to the following:
220 | a) Identity of the reagent or consumable? | Yes
221 | b) Batch code or lot number? | Yes
222 | c) Manufacturer or supplier name and contact information? | Yes
223 | d) Date of receiving, the expiry date, date of entering into service and, where applicable, the date the material was taken out of service? | Yes
224 | e) Manufacturer's instruction/package insert? | Yes
225 | f) Records of inspection of reagents and consumables when received (e.g. acceptable or damaged)? | Yes
226 | 7.5 Budgetary Projections | Yes
227 | 7.6 Management Review of Supply Requests | Yes
228 | 7.7 Laboratory Inventory System
229 | a) Are inventory records complete and accurate, with minimum and maximum stock levels denoted and monitored? | Yes
230 | b) Is the consumption rate of all reagents and consumables monitored? | Yes
231 | c) Are stock counts routinely performed? | Yes
232 | 7.8 Storage AreaAre storage areas set up and monitored appropriately?
233 | a) Is the storage area well-organized and free of clutter? | Yes
234 | b) Are there designated places labeled for all inventory items for easy access? | Yes
235 | c) Is adequate cold storage available? | Yes
236 | d) Are storage areas monitored as per prescribed storage conditions? | Yes
237 | e) Is the ambient temperature monitored routinely? | Partial | not done regularly
238 | f) Is storage in direct sunlight avoided? | Yes
239 | g) Is the storage area adequately ventilated? | Yes
240 | h) Is the storage area clean and free of dust and pests? | Yes
241 | i) Are storage areas access-controlled? | Yes
242 | 7.9 Inventory Organization and Wastage Minimization | Yes
243 | 7.10 Disposal of Expired Products | Yes
244 | 7.11 Product Expiration | No | no tracking
245 | 7.12 Laboratory Testing Services | Yes
246 | 8.1 Information for patients and users | Yes
247 | 8.2 Does the laboratory adequately collect information needed for examination performance?
248 | a) Are all test requests accompanied by an acceptable and approved test requisition form (and a transmittal sheet/checklist/manifest where applicable)? | Yes
249 | b) Does the request form has patient ID including gender, date of birth, location of patient and unique identifier? | Yes
250 | c) Name, signature or initials of authorized requester | Yes
251 | d) Type of sample and examination requested | Yes
252 | e) Clinically relevant information | Yes
253 | f) Date of sample collection (And time of collection where relevant – where time has an impact on the result) | Yes
254 | g) Date and time of sample receipt | Yes
255 | h) Written consent for invasive procedures with increased risk of complications | Yes
256 | 8.3 Are adequate sample receiving procedures in place?
257 | a) Patient Unique Identifier | Yes
258 | b) Are received specimens evaluated according to acceptance/rejection criteria? | Yes
259 | c) Are specimens logged appropriately upon receipt in the laboratory (including date, time, and name of receiving officer)? | No | Receiving officer missing
260 | d) Are procedures in place to process 'urgent' specimens and verbal requests? | No | No procedure
261 | e) When samples are split, can the portions be traced back to the primary sample? | Yes
262 | f) If not a 24 hour lab, is there a documented method for handling of specimens received after hours? | Yes
263 | g) Are specimens delivered to the correct workstations in a timely manner? | Yes
264 | 8.4 Pre-examination Handling, Preparation and Storage | Yes
265 | 8.5 Sample Transportation | Yes
266 | 8.6 Does the laboratory select and evaluate referral Labs and Consultants?
267 | a) Are there documented reviews and evaluations of referral laboratories and consultants as defined by the laboratory? | Yes
268 | b) Is there a register of referral Laboratories and consultants? | Yes
269 | c) Are referred specimens tracked properly using a logbook, tracking form or electronically? | Yes
270 | 8.7 Documentation of Examination Procedures | Yes
271 | 8.8 Reagents Acceptance Testing | Yes
272 | 8.9 Quality Control | Partial | No QC on sickling test
273 | 8.10 Quality Control DataAre QC results monitored and reviewed (including biases and Levy-Jennings charts for quantitative tests)?
274 | a) Is there documentation of corrective action taken when quality control results exceed the acceptable range or reviews identify non conformities in a timely manner? | Yes
275 | b) Does the Lab evaluate the results from the patient samples that were examined after the last successful quality control event | Yes
276 | 8.11 Comparability of Examination ResultsDoes the laboratory compare results of the same test performed with different procedures and equipment?
277 | a) Where there is more than one procedure for the same measure, does the laboratory compare results from the different procedures, equipment or methods? | Yes
278 | b) Does the lab discuss, document and act upon (including notifying users) problems or deficiencies from these comparison studies? | Yes
279 | 8.12 Are environmental conditions checked and reviewed accurately?8.12 Are the following environmental conditions checked and recorded daily?
280 | a) Room temperature | Yes
281 | b) Freezers | Yes
282 | c) Refrigerator | Yes
283 | d) Incubators | Yes
284 | e) Water Bath | Yes
285 | 8.13 Have acceptable ranges been defined for all temperature- dependent equipment with procedures and documentation of action taken in response to out of range temperatures? | Yes
286 | 8.14 Does the laboratory participate in interlaboratory comparison program or alternative assessment systems for all tests?
287 | a) Do samples come from providers who are accredited or approved? | Yes
288 | b) Are specimens handled and tested the same way as patient specimens? | Yes
289 | c) Is the performance of the laboratory in the PT program reviewed and discussed with relevant staff? | Partial | Not all PT are reviewed (no documentary evidence of review Malaria PT of March 2017)
290 | d) Is cause analysis performed for unacceptable results? | Yes
291 | e) Is corrective action documented for unacceptable results? | Yes
292 | 9.1 Test Result Reporting System | Yes
293 | 9.2 Testing Personnel | Yes
294 | 9.3 Report ContentDoes the laboratory report contain at least the following:
295 | a) Test requested | Yes
296 | b) Identification of the laboratory | Yes
297 | c) Identification of all examinations performed by a referral laboratory | Yes
298 | d) Patient identification and location | Yes
299 | e) Name of the requester | Yes
300 | f) Date of primary sample collection (and time, relevant to patient care) | Yes
301 | g) Type of primary sample | Yes
302 |  | Yes
303 | i) Biological reference intervals where applicable | Yes
304 | j) Is there space for interpretation or comments of results, when applicable? | Yes
305 | k) Identification of the person(s) reviewing and authorizing the report | Yes
306 | l) Date and time of the report | Partial | not consistently
307 | m) Page number to total number of pages (e.g. 'Page 1 of 5', 'Page 2 of 5', etc.) | No | No paging
308 | n) When issuing revised reports, is it clearly identified as a revision and includes reference to the date and patient's identity in the original report and the user made aware of the revision? | Yes
309 | o) Does the revised record show the time and date of the change and the name of the person responsible for the change? | Yes
310 | p) Does the original report entry remain in the record when revisions are made? | Yes
311 | 9.4 Analytic System/Method Tracing | Yes
312 | 9.5 Archived Data Labeling and Storage | Yes
313 | 9.6 Authorities and Responsibilities Has the laboratory defined and implemented authorities and responsibilities for the management and use of the laboratory information system– paper based and electronic, including maintenance and modifications that may affect patient care? Is the following in place and implemented?
314 | a) Controlled access to patient data and information | Yes
315 | b) Controlled access to enter patient data and examination results | Yes
316 | c) Controlled access to changing patient data or examination results | Yes
317 | d) Controlled access to the release of examination results and reports | Yes
318 | e) Verify that results that have been transmitted electronically or reproduced external to the laboratory (computers, fax machines, email and websites and personal web devices) are correct | Yes
319 | 9.7 Information Management System | Yes
320 | 9.8 Test Result Report | Yes
321 | 9.9 Verification of Electronic Laboratory Information System
322 | a) Has the system been verified before implementation that include the verification reports to check functioning and inter-phasing by the laboratory?
323 | b) Records of the validation by the supplier available and approved for use?
324 | c) Ongoing system checks available for correct transmissions, calculations and storage of results and records.
325 | 9.10 Is the Laboratory Information System properly maintained to ensure continued functioning:
326 | a) Documented regular service by authorized and trained personnel | Yes
327 | b) Documented system failures with documented appropriate root cause analysis, corrective actions and preventative actions | Not Applicable
328 | c) System operated in an environment recommended by the supplier for optimal functioning | Yes
329 | 10.1 Are all identified nonconforming activities/ work identified and documented adequately?
330 | a) Indicating details of what happened, when, person responsible. | Yes
331 | b) Immediate actions being taken | Yes
332 | c) Determination of the extent of the non- conformity? | Yes
333 | d) Are examinations halted and results withheld or recalled where the non-conformity compromises patient results? | Yes
334 | e) Informing the requester where the non -conformity has an effect on the management of the patient | Yes
335 | f) Authorization of resumption of testing documented (where testing has been halted) | Yes
336 | 10.2 Is documented root cause analysis performed for non-conforming work before corrective actions are implemented? | Yes
337 | 10.3 Is corrective action performed and documented for non-conforming work? | Yes
338 | 10.4 Are implemented corrective actions monitored and reviewed for their effectiveness before closure/clearance? | Yes
339 | 10.5 Preventive ActionsAre documented preventive actions implemented and monitored for their effectiveness?
340 | a) Reviewing of laboratory data and information to determine potential non conformities | Yes
341 | b) Determining root causes for potential non conformities | Yes
342 | c) Implementing and documenting preventive actions | Yes
343 | d) Reviewing and documenting effectiveness of preventive actions | Yes
344 | 11.1 Graphical Tools | Yes
345 | 11.2 Quality Management System Improvement Measures | Yes
346 | 11.3 Communication System on Laboratory Operations | Yes
347 | 11.4 Quality Indicators | Yes
348 | 11.5 Lab Performance Improvement | Yes
349 | 11.6 Quality of Lab Performance | Yes
350 | 12.1 Size of the Laboratory | Yes
351 | 12.2 Patient care and testing areas | No | not seperated
352 | 12.3 Individual workstation maintenanceIs each individual workstation maintained free of clutter and set up for efficient operation?
353 | a) Does the equipment placement/layout facilitate optimum workflow? | Yes
354 | b) Are all needed supplies present and easily accessible? | Yes
355 | c) Are the chairs/stools at the workstations appropriate for bench height and the testing operations being performed? | Yes
356 | 12.4 Physical work environmentIs the physical work environment appropriate for testing?
357 | a) Free of clutter | Yes
358 | b) Adequately ventilated? | Yes
359 | c) Adequately lit? | Yes
360 | d) Climate-controlled for optimum equipment function? | Yes
361 | e) Are filters checked, cleaned and/or replaced at regular intervals, where air-conditioning is installed? | Yes
362 | f) Are wires and cables properly located and protected from traffic? | Yes
363 | g) Is there a functioning back-up power supply (generator)? | Yes
364 | h) Is critical equipment supported by uninterrupted power source (UPS) systems? | Yes
365 | i) Is equipment placed appropriately (away from water hazards, out of traffic areas)? | Yes
366 | j) Are appropriate provisions made for adequate water supply, including deionized water (DI) or distilled water, if needed? | Yes
367 | k) Is clerical work completed outside the testing area? | Yes
368 | l) Is major safety signage posted and enforced including NO EATING, SMOKING, DRINKING? | Yes
369 | 12.5 Laboratory Access | Yes
370 | 12.6 Laboratory Storage Areas | Yes
371 | 12.7 Work Area | Yes
372 | 12.8 Biosafety Cabinet | Yes
373 | 12.9 Laboratory Safety ManualIs a laboratory safety manual available, accessible, and up-to-date? Does the safety manual include guidelines on the following topics?
374 | a) Blood and Body Fluid Precautions | Yes
375 | b) Hazardous Waste Disposal | Yes
376 | c) Hazardous Chemicals / Materials | Yes
377 | d) MSDS Sheets | Yes
378 | e) Personal protective equipment | Yes
379 | f) Vaccination | Yes
380 | g) Post-Exposure Prophylaxis | Yes
381 | h) Fire Safety | Yes
382 | i) Electrical safety | Yes
383 | 12.10 Waste Disposal | Yes
384 | 12.11 Hazardous ChemicalsAre hazardous chemicals / materials properly handled?
385 | a) Are hazardous chemicals properly labeled? | Yes
386 | b) Are hazardous chemicals properly stored? | Yes
387 | c) Are hazardous chemicals properly utilized? | Yes
388 | d) Are hazardous chemicals properly disposed? | Yes
389 | 12.12 Handling of Sharps | Yes
390 | 12.13 Fire SafetyIs fire safety included as part of the laboratory’s overall safety program?
391 | a) Are all electrical cords, plugs, and receptacles used appropriately and in good repair? | Yes
392 | b) Is an appropriate fire extinguisher available, properly placed, in working condition, and routinely inspected? | Yes
393 | c) Is an operational fire warning system in place? | Yes
394 | d) Are periodic fire drills conducted at defined period of time? | Yes
395 | 12.14 Safety AuditsAre safety inspections or audits conducted regularly and documented?
396 | a) Is there an audit plan/schedule that ensures all activities of the lab are checked for safety compliance? | Yes
397 | b) Are inspections/audits being carried out by authorized persons? | Yes
398 | c) Are the personnel conducting the internal audits trained in safety? | Yes
399 | d) Is cause analysis and action taken for nonconformities/noted deficiencies? | Yes
400 | e) Are safety findings documented and presented to the laboratory management and relevant staff for review? | Yes
401 | 12.15 Safety EquipmentIs standard safety equipment available and in use in the laboratory?
402 | a) Biosafety cabinet(s) | Yes
403 | b) Covers on centrifuge(s) | Yes
404 | c) Hand-washing station | Yes
405 | d) Eyewash station/bottle(s) and showers where applicable | Yes
406 | e) Spill kit(s) | No | no found
407 | f) First aid kit(s) | Yes
408 | 12.16 Personnel Protective Equipment | Yes
409 | 12.17 Staff Vaccinations | Yes | not all staff
410 | 12.18 Post Exposure Prophylaxis | Yes
411 | 12.19 Adverse incidents | Yes
412 | 12.20 Biosafety Training | Yes
413 | 12.21 Laboratory Safety Officer | Yes

## Scores
Question | Description | Points
1 | 1.1 | 2
2 | 1.2 | 1
3 | 1.3 | 2
4 | 1.4 | 2
5 | 1.5 | 1
6 | 1.6 | 2
7 | 1.7 | 2
8 | 1.8 | 2
9 | 1.9 | 2
10 | 1.1 | 2
11 | 1.11 | 2
12 | 2.1 | 1
13 | 2.2 | 1
14 | 2.3 | 2
15 | 2.4 | 2
16 | 3.1 | 2
17 | 3.2 | 2
18 | 3.3 | 1
19 | 3.4 | 3
20 | 3.5 | 1
21 | 3.6 | 3
22 | 3.7 | 1
23 | 3.8 | 1
24 | 4.1 | 2
25 | 4.2 | 2
26 | 4.3 | 2
27 | 4.4 | 2
28 | 4.5 | 0
29 | 6.1 | 1
30 | 6.2 | 5
31 | 6.3 | 5
32 | 7.1 | 2
33 | 7.2 | 1
34 | 7.3 | 2
35 | 7.4 | 2
36 | 7.5 | 2
37 | 7.6 | 2
38 | 7.7 | 2
39 | 7.8 | 1
40 | 7.9 | 2
41 | 7.1 | 2
42 | 7.11 | 0
43 | 7.12 | 2
44 | 10.1 | 5
45 | 10.2 | 3
46 | 10.3 | 3
47 | 10.4 | 3
48 | 10.5 | 5
49 | 11.1 | 2
50 | 11.2 | 2
51 | 11.3 | 2
52 | 11.4 | 2
53 | 11.5 | 2
54 | 11.6 | 2
55 | 5.1 | 2
56 | 5.2 | 2
57 | 5.3 | 5
58 | 5.4 | 1
59 | 5.5 | 2
60 | 5.6 | 2
61 | 5.7 | 0
62 | 5.8 | 2
63 | 5.9 | 2
64 | 5.1 | 2
65 | 5.11 | 2
66 | 5.12 | 2
67 | 5.13 | 2
68 | 5.14 | 2
69 | 5.15 | 0
70 | 5.16 | 2
71 | 12.1 | 2
72 | 12.2 | 0
73 | 12.3 | 2
74 | 12.4 | 2
75 | 12.5 | 2
76 | 12.6 | 2
77 | 12.7 | 2
78 | 12.8 | 2
79 | 12.9 | 2
80 | 12.1 | 2
81 | 12.1 | 2
82 | 12.1 | 2
83 | 12.1 | 2
84 | 12.1 | 3
85 | 12.1 | 1
86 | 12.1 | 2
87 | 12.1 | 2
88 | 12.1 | 2
89 | 12.1 | 2
90 | 12.2 | 2
91 | 12.2 | 2
92 | 9.1 | 2
93 | 9.2 | 2
94 | 9.3 | 1
95 | 9.4 | 2
96 | 9.5 | 2
97 | 9.6 | 2
98 | 9.7 | 2
99 | 9.8 | 2
100 | 9.1 | 2
101 | 8.1 | 2
102 | 8.2 | 3
103 | 8.3 | 1
104 | 8.4 | 2
105 | 8.5 | 2
106 | 8.6 | 2
107 | 8.7 | 2
108 | 8.8 | 2
109 | 8.9 | 1
110 | 8.1 | 3
111 | 8.11 | 2
112 | 8.12 | 2
113 | 8.13 | 2
114 | 8.14 | 1

## Summary of Assessment Findings
Field | Value
Commendations | 1.0 DOCUMENTS AND RECORDS b): 232 f): to to to Communication (internal and external): due to this 2.0 MANAGEMENT REVIEWS AND MANAGEMENT RESPONSIBILITIES b): No action items d): Not reviewed consistently l): Not documented f): Not used i): not done consistently Review Output: Does the management review meeting include the following outputs?: ddddhu 3.0 ORGANIZATION AND PERSONNEL a): no budgets f): not done m): No records of medical surveillance b): not documented g): no feedback i): nor done 4.0 CLIENT MANAGEMENT AND CUSTOMER SERVICE No evaluation tool 5.0 EQUIPMENT b): Not documented c): Not documented Broken down centrifuge not labeled Water bath manual missing 6.0 EVALUATIONS AND AUDITS c): no certs 7.0 PURCHASING AND INVENTORY CONTROL only one evaluated e): not done regularly no tracking 8.0 PROCESS CONTROL c): Receiving officer missing d): No procedure No QC on sickling test c): Not all PT are reviewed (no documentary evidence of review Malaria PT of March 2017) 9.0 INFORMATION MANAGEMENT l): not consistently m): No paging 12.0 FACILITIES & BIOSAFETY not seperated e): no found not all staff
Challenges | no challenges encountered
Recommendations | accreditation recommended up on closure of the non- conformance identified

## Action Plan
Action | Incharge | Timeline