## Laboratory Information
Field | Value
Laboratory Name | ASPE Medical Clinic
Laboratory Number | 0023
Laboratory Address | P.O. Box 59857-00100, Nairobi
Laboratory Telephone | 0703034000
Laboratory Fax | 6007498
Laboratory Email | aspe@aspe.org
Head of Laboratory | Kiprono Chepkonga
Telephone (Head of Laboratory - Personal) | 0722222222
Telephone (Head of Laboratory - Work) | 0713492839
Lab Level | National
Lab Affiliation | Private

## Staffing Summary
Profession | Employees | Adequate
Degree-holding Professional Staff | 13 | No
Diploma-holding Professional Staff | 17 | Yes
Certificate-holding Professional Staff | 1 | No
Microscopist |  | Insufficient Data
Data Clerk | 1 | No
Phlebotomist |  | Insufficient Data
Cleaner | 1 | No
Is the cleaner(s) dedicated to the laboratory only? | Yes
Has the cleaner(s) been trained in safe waste handling? | No
Driver | 1 | Insufficient Data
Is the driver(s) dedicated to the laboratory only? | No
Has the driver(s) been trained in biosafety? | No
Other |  | Insufficient Data

## Organizational Structure
Field | Value
Sufficient Space | Yes
Equipment | Yes
Supplies | Yes
Personnel | Yes
Infrastructure | Yes
Other - Please specify:: | Insufficient Data

## SLMTA Information
Field | Value
Official SLMTA Audit? | Yes
Audit Start Date | 2015-04-07
Audit End Date | 2015-04-14
Name(s) and Affiliation(s) of Auditor(s) | Kitsao Emmanuel
Type of SLMTA Audit | Midterm Audit
Total # of tests before SLMTA | 21167
Total # of tests this year | 591
Cohort ID | 19
SLMTA Labtype | Non-Governmental Organization
Date of baseline audit | 2014-03-30
Date of SLMTA workshop #1 | 2014-04-21
Date of exit audit performed | 2015-03-31
Baseline Score | 231
Stars obtained | 3 Stars
Exit Score | 240
Stars obtained | 3 Stars
Date of Last Audit | 2015-03-31
Score for Last Audit | 243
Prior Audit Status | 3 Stars

## Assessment Details
Question | Description | Response | Notes | Compliance
1 | 1.1 Laboratory Quality ManualIs there a current laboratory quality manual, composed of the quality management system’s policies and procedures, and has the manual content been communicated to and understood and implemented by all staff?
2 | 1.1.1 Structure defined per ISO15189, Section 4.2.4 | Yes
3 | 1.1.2 Quality policy statement that includes scope of service, standard of service, objectives of the quality management system, and management commitment to compliance | Yes
4 | 1.1.3 Description of the quality management system and the structure of its documentation | Yes
5 | 1.1.4 Reference to supporting procedures, including technical procedures | Yes
6 | 1.1.5 Description of the roles and responsibilities of the laboratory manager, quality manager, and other personnel responsible for ensuring compliance | Yes
7 | 1.1.6 Documentation of at least annual management review and approval. | Yes
8 | 1.2 Document and Information Control System | Partial | Biosafety/Biosecurity guidance not controlled
9 | 1.3 Document and Records | Partial | Master list has some non-existent documents
10 | 1.4 Laboratory Policies and Standard Operating ProceduresAre policies and standard operating procedures (SOPs) for laboratory functions current, available and approved by authorized personnel?
11 | 1.4.1 Document & Record Control | Yes | GEN/SOP/016
12 | 1.4.2 Conflict of Interest | Yes | GEN/SOP/013
13 | 1.4.3 Communication | Yes | GEN/SOP/018
14 | 1.4.4 Review of Contracts (Supplier and Customer) | Yes | GEN/SOP/019
15 | 1.4.5 Examination by Referral Laboratories | Yes | GEN/SOP/004
16 | 1.4.6 Purchasing and Inventory Control | Yes
17 | 1.4.7 Advisory Services | Yes | GEN/SOP/021
18 | 1.4.8 Resolution of Complaints and Feedback | Yes | GEN/SOP/007
19 | 1.4.9 Identification and Control of Nonconformities | Yes | GEN/SOP/008
20 | 1.4.10 Corrective Action | Yes | GEN/SOP/008
21 | 1.4.11 Preventive Action | Yes | GEN/SOP/008
22 | 1.4.12 Continual Improvement | Yes | GEN/SOP/008
23 | 1.4.13 Quality and Technical Records | Yes | GEN/SOP/002
24 | 1.4.14 Internal Audits | Yes | GEN/SOP/009
25 | 1.4.15 Management Review | Yes | GEN/SOP/010
26 | 1.4.16 Personnel Records/Files | Yes | GEN/SOP/008
27 | 1.4.17 Personnel Training | No | No definition of appraisals
28 | 1.4.18 Competency audit | No | No pass/fail criteria
29 | 1.4.19 Authorization | Yes | GEN/SOP/008
30 | 1.4.20 Accommodation and Environmental Conditions | Yes | GEN/SOP/014
31 | 1.4.21 Equipment | No | In progress
32 | 1.4.22 Calibration of Equipment | No | In progress
33 | 1.4.23 Pre-examination Procedures (Handbook) | Yes
34 | 1.4.24 Specimen Storage and Retention | No
35 | 1.4.25 Examination SOPs | Yes
36 | 1.4.26 Equipment Validation/Verification | Yes
37 | 1.4.27 Interrupted Services | No
38 | 1.4.28 Examination Validation/Verification | No
39 | 1.4.29 Quality Assurance | Yes
40 | 1.4.30 Reporting of Results | Yes
41 | 1.4.31 Patient Confidentiality | Yes | Quality manual 5.7 GEN SOP/006
42 | 1.4.32 Laboratory Safety or Safety Manual | Yes
43 | 1.5 Policy and SOPs Accessibility | Yes
44 | 1.6 Policies and SOPs Communication | Partial | Some SOPs not read by all staff
45 | 1.7 Document Control Log | Partial | Superseded quality manual does not indicate date of discontinuation
46 | 1.8 Discontinued Policies and SOPs | Yes
47 | 1.9 Data Files | Yes
48 | 1.10 Archived Results Accessibility | No | Records not arranged in organized manner. No list of archived records
49 | 2.1 Workplan and Budget | Yes
50 | 2.2 Review of Quality and Technical RecordsDoes the laboratory supervisor routinely perform a documented review of all quality and technical records?
51 | 2.2.1 Follow-up of action items from previous reviews | Yes
52 | 2.2.2 Status of corrective actions taken and required preventive actions | Yes
53 | 2.2.3 Reports from personnel | Yes
54 | 2.2.4 Changes in volume and type of work the laboratory undertakes | Yes
55 | 2.2.5 Changes in the suitability of biological reference ranges | Yes
56 | 2.2.6 Changes in the client handbook | Yes
57 | 2.2.7 Environmental monitoring log sheets | No
58 | 2.2.8 Specimen rejection logbook | No
59 | 2.2.9 Equipment calibration and maintenance records | No
60 | 2.2.10 IQC records across all test areas | No
61 | 2.2.11 Outcomes of PTs and other forms of Inter-laboratory comparisons | Yes
62 | 2.2.12 Monitoring of turnaround time | Yes
63 | 2.2.13 Quality indicators | Yes
64 | 2.2.14 Outcomes from recent internal audit records | Yes
65 | 2.2.15 Results of audit(s) or audits by external bodies | No
66 | 2.2.16 Customer complaints and feedback | Yes
67 | 2.2.17 Occurrence/incidence logs, nonconformities and corrective action reports | Yes
68 | 2.2.18 Results of improvement projects | Yes
69 | 2.2.19 Operational procedures (for potential sources of non-conformance and opportunities for improvement) | Yes
70 | 2.2.20 Evaluation of performance of referral laboratories | Yes
71 | 2.2.21 Evaluation of supplier performance | Yes
72 | 2.2.22 Document Review | Yes
73 | 2.2.23 Documentation of review and action planning with staff for resolution and follow-up review | Yes
74 | 2.3 Annual Review of Quality Management SystemsDoes the laboratory management annually perform a review of all quality systems at a management review meeting?
75 | 2.3.1 Follow-up of action items from previous management reviews | Yes
76 | 2.3.2 Status of corrective actions taken and required preventive actions | Yes
77 | 2.3.3 Reports from managerial and supervisory personnel | Yes
78 | 2.3.4 Changes in volume and type of work the laboratory undertakes | No
79 | 2.3.5 Changes in the suitability of biological reference ranges | No
80 | 2.3.6 Changes in the client handbook | Yes
81 | 2.3.7 Environmental monitoring log sheets | Yes
82 | 2.3.8 Specimen rejection logbook | Yes
83 | 2.3.9 Equipment calibration and maintenance records | No
84 | 2.3.10 IQC records across all test areas | No
85 | 2.3.11 Outcomes of PTs and other forms of Inter-laboratory comparisons | No
86 | 2.3.12 Turnaround time | No
87 | 2.3.13 Quality indicators | Yes
88 | 2.3.14 Outcomes from recent internal audit records | Yes
89 | 2.3.15 Results of audit(s) or audits by external bodies | Yes
90 | 2.3.16 Customer complaints and feedback | Yes
91 | 2.3.17 Occurrence/incidence logs, nonconformities and corrective action reports | Yes
92 | 2.3.18 Results of improvement projects | Yes
93 | 2.3.19 Operational procedures (for potential sources of non-conformance and opportunities for improvement) | Yes
94 | 2.3.20 Evaluation of performance of referral laboratories | Yes
95 | 2.3.21 Evaluation of supplier performance | Yes
96 | 2.3.22 Document Review | Yes
97 | 2.3.23 Documentation of review and action planning with staff for resolution and follow-up review | No
98 | 2.4 Quality Management System Improvement Measures | Yes | TAT, Sample rejection
99 | 2.5 Communications System on Laboratory Operations | Yes | Genexpert discussed in HMT
100 | 3.1 Workload, Schedule and Coverage | Yes
101 | 3.2 Duty Roster And Daily Routine | Yes
102 | 3.3 Organizational Chart and External/Internal Reporting Systems | Yes
103 | 3.4 Quality Management System Oversight | Yes
104 | 3.5 Personnel Filing SystemAre Personnel Files present?
105 | 3.5.1 Employee Orientation | Yes | No available for Karisa
106 | 3.5.2 Education & Training (e.g., degrees/certificates) | Yes
107 | 3.5.3 Previous experience and work history (e.g. CV) | Yes
108 | 3.5.4 Written job description with documentation that staff member received and signed a copy of their job description | Yes
109 | 3.5.5 Letter of employment or appointment | Yes
110 | 3.5.6 Review of job-relevant SOPs | Yes | SOP manual log
111 | 3.5.7 Documented review of safety manual, evidence of safety training | Yes | Safety manual
112 | 3.5.8 Review of procedure for employees to communicate concerns about test quality and laboratory safety | Yes | Staff meetings
113 | 3.5.9 Registration with professional board | Yes
114 | 3.5.10 Training record documenting trainings received, vendor training received on-site | Yes
115 | 3.5.11 Periodic Performance Review – including Observation, Competency audit, Coaching / Feedback, on-the-job training | Yes | Objective evidence not provided
116 | 3.5.12 Documentation of employee recognition (i.e., employee of the month, letter of commendation, etc.) | Yes | Hospital program
117 | 3.5.13 Human Resource (HR) Data – (vaccination status, accidental exposure during work injuries, accident history, leave days taken, etc.) | Yes
118 | 3.6 Staff Competency audit and Training | Partial | No pass/fail criteria. No raw data for competency assessment
119 | 3.7 Laboratory Staff Training | Partial | No training plans. CME schedules
120 | 3.8 Staff MeetingsAre staff meetings held regularly?
121 | 3.8.1 Follow-up of action items from previous staff meetings | Yes | 20-03-2013
122 | 3.8.2 Discussion about problems and complaints | Yes
123 | 3.8.3 Review of documentation | Yes | 26-07-2013
124 | 3.8.4 Communication on reviewed/revised/redundant SOPs | Yes
125 | 3.8.5 Systemic and or recurrent problems and issues addressed, including actions to prevent recurrence | Yes | 23-07-2013
126 | 3.8.6 Review of results from prior corrective actions | Yes
127 | 3.8.7 Discussion and evaluation of improvement topics/projects | No
128 | 3.8.8 Feedback given by staff that have attended meetings, training, conferences etc. | No
129 | 3.8.9 Recognition of employees for exemplary performance (i.e., employee of the month, letter of commendation, etc.) | No
130 | 3.8.10 Relay of reports and updates from lab attendance at meetings with clinicians (the use of lab services and/or attendance at clinical rounds) | No
131 | 3.8.11 Recording and monitoring of meeting notes for progress on issues | Yes
132 | 4.1 Advice and Training by Qualified Staff | Yes
133 | 4.2 Laboratory Handbook for Clients | Yes | CPGH/MAN/CHB/0032
134 | 4.3 Communication Policy on Delays in Service | Yes | Interruption notices available
135 | 4.4 Evaluation Tool and Follow up | Partial | Customer survey report on July 2013. Not reviewed by management - actions not implemented
136 | 5.1 Adherence to Proper Equipment Protocol | Yes
137 | 5.2 Equipment and Method Validation/Verification and Documentation | Partial | Validation has not been done for the Chemistry Cobas C111, Celtac F and Facs Calibur
138 | 5.3 Equipment Record MaintenanceIs current equipment inventory data available on all equipment in the laboratory?
139 | 5.3.1 Name of equipment | Yes
140 | 5.3.2 Manufacturer’s contact details | Yes
141 | 5.3.3 Condition received (new, used, reconditioned) | No
142 | 5.3.4 Serial number | No | Serial number for Facs Calibur
143 | 5.3.5 Date of purchase | No | Not indicated
144 | 5.3.6 Date when put 'out of service' | No | Not indicated
145 | 5.3.7 Date of entry into service | No | Missing
146 | 5.4 Equipment Maintenance RecordsIs relevant equipment service information readily available in the laboratory?
147 | 5.4.1 Service contract information | Yes
148 | 5.4.2 Contact details for service provider | Yes
149 | 5.4.3 Decontamination Records | No
150 | 5.4.4 Performance and maintenance records | No
151 | 5.4.5 Last date of service | Yes
152 | 5.4.6 Next date of service | Yes
153 | 5.4.7 Current location | No | Not indicated
154 | 5.5 Obsolete Equipment Procedures | Yes
155 | 5.6 Adherence to Equipment Calibration Protocol | Partial | Pipettes are due for calibration
156 | 5.7 Equipment Preventive Maintenance | Partial | Operator has never done any maintenance on the centrifuge
157 | 5.8 Equipment Service Maintenance | Partial | No schedule for maintenance of fridges, incubator and centrifuge
158 | 5.9 Equipment Parts for Repair | Yes
159 | 5.10 Equipment Malfunction - Response and Documentation | Partial | No corrective action for breakdowns e.g. Facs Calibur
160 | 5.11 Equipment Repair Monitoring and Documentation | Partial | No evidence for documentation of QC or calibration after service or repair
161 | 5.12 Equipment Failure - Contingency Plan | Partial | No documented procedure
162 | 5.13 Manufacturer's Operator Manual | Yes
163 | 5.14 Communication on Effectiveness of Quality Management System | Yes
164 | 5.15 Laboratory Testing Services | Partial | Service interrupted due to failure of Cobas C111
165 | 6.1 Internal AuditsAre internal audits conducted at intervals as defined in the quality manual and do these audits address areas important to patient care?
166 | 6.1.1 Are audits being carried out by persons who are not involved in lab activities in the section being audited? | Yes
167 | 6.1.2 Are the personnel conducting the internal audits trained and competent in auditing? | No | One of the staff not trained on the procedures
168 | 6.1.3 Is cause analysis performed for nonconformities/noted deficiencies? | No | Some NCs not closed
169 | 6.1.4 Are internal audit findings documented and presented to the laboratory management and relevant staff for review? | No
170 | 6.2 Audit Recommendations and Action Plan & Follow up | Partial | Congestion in microbiology still outstanding
171 | 7.1 Inventory and Budgeting System | Yes
172 | 7.2 Service Supplier Performance Review | Yes
173 | 7.3 Manufacturer/Supplier List | Yes
174 | 7.4 Budgetary Projections | Yes
175 | 7.5 Management Review of Supply Requests | Yes
176 | 7.6 Order Tracking, Inspection, and Documentation | Partial | Facs flow reagent were not marked with receiving dates
177 | 7.7 Inventory Control SystemIs an inventory control system in place?
178 | 7.7.1 Acceptance and rejection of consumables | Yes
179 | 7.7.2 Recording of lot number, date of receipt, received by and date placed into service | No | Lot number not recorded for Tri test
180 | 7.7.3 Storage of consumables | Yes
181 | 7.8 Laboratory Inventory System | Yes
182 | 7.9 Usage Rate Tracking of Consumables | Yes
183 | 7.10 Inventory Control System – Stock Counts | Yes
184 | 7.11 Storage AreaAre storage areas set up and monitored appropriately?
185 | 7.11.1 Is the storage area well-organized and free of clutter? | Yes
186 | 7.11.2 Are there designated places labeled for all inventory items? | Yes
187 | 7.11.3 Are hazardous chemicals stored appropriately? | Yes
188 | 7.11.4 Is adequate cold storage available? | Yes
189 | 7.11.5 Are storage areas monitored as per prescribed storage conditions? | Yes
190 | 7.11.6 Is the ambient temperature monitored routinely? | Yes
191 | 7.11.7 Is storage in direct sunlight avoided? | Yes
192 | 7.11.8 Is the storage area adequately ventilated? | Yes
193 | 7.11.9 Is the storage area clean and free of dust and pests? | Yes
194 | 7.11.10 Are storage areas access-controlled? | Yes
195 | 7.12 Inventory Organization and Wastage Minimization | Yes
196 | 7.13 Disposal of Expired Products | Partial | Xylene directly disposed to sink contrary to MSDS
197 | 7.14 Product Expiration | No | Salmonella typhi reagent expired
198 | 7.15 Laboratory Testing Services | Partial | CD4 testing was interrupted due to stock out
199 | 8.1 Are guidelines for patient identification, specimen collection (including client safety), labeling, and transport readily available to persons responsible for primary sample collection? | Yes
200 | 8.2 Are adequate sample receiving procedures in place?
201 | 8.2.1 Are specimens labeled with patient ID, test, and date, time of collection, date of collection and authorized requester? | No | Request missing date, requester and time e.g. 156817
202 | 8.2.2 Are all test requests accompanied by an acceptable and approved test requisition form? | No | Request done on small pieces of paper
203 | 8.2.3 If not a 24 hour lab, is there a documented method for handling of specimens received after hours? | Yes
204 | 8.2.4 Are all samples that are either received or referred to a higher level laboratory accompanied by a sample delivery checklist or transmittal sheet? | Yes
205 | 8.2.5 Are received specimens evaluated according to acceptance/rejection criteria? | Yes
206 | 8.2.6 Are specimens logged appropriately upon receipt in the laboratory (including date, time, and name of receiving officer)? | No | Casualty la logbook missing date and tome
207 | 8.2.7 When samples are split, can the portions be traced back to the primary sample? | Yes
208 | 8.2.8 Is a two-identifier system in use and is each sample assigned a unique identifying number? | Yes
209 | 8.2.9 Are procedures in place to process 'urgent' specimens and verbal requests? | No | No procedure for urgent tests
210 | 8.2.10 Are specimens delivered to the correct workstations in a timely manner? | Yes
211 | 8.3 Are specimens stored appropriately prior to testing?
212 | 8.3.1 Are specimens disposed of in a safe manner? | Yes
213 | 8.4 Are specimens packaged appropriately according to local and or international regulations and transported to referral laboratories within acceptable timeframes? | Yes
214 | 8.5 Are referred specimens tracked properly using a logbook or tracking form? | Yes
215 | 8.6 Is complete procedure manual available at the workstation or in the work area? | Yes
216 | 8.7 Is there a reagent logbook for lot number and dates of opening that reflects verification of new lots? | Yes
217 | 8.8 Is each new lot number, new shipment of reagents, or consumables verified before use? | Partial | No lot verification
218 | 8.9 Is internal quality control performed, documented, and verified before releasing patient results? | Partial | In Chemistry lab, QC not done
219 | 8.10 Are QC results monitored and reviewed (biases, shifts, trends, and Levy-Jennings charts)? Is there documentation of corrective action when quality control results exceed the acceptable range in a timely manner? | Yes
220 | 8.11 Are environmental conditions are checked and reviewed accurately?
221 | 8.11.1 Room temperature | Yes
222 | 8.11.2 Freezers | Yes
223 | 8.11.3 Refrigerator | Yes
224 | 8.11.4 Incubators | Yes
225 | 8.11.5 Water Bath | Yes
226 | 8.12 Have acceptable ranges been defined for all temperature- dependent equipment with procedures and documentation of action taken in response to out of range temperatures? | Yes
227 | 8.13 Does the laboratory participate in external Proficiency Testing (PT) or exercise an alternative performance audit system when appropriate?
228 | 8.13.1 Are blinded characterized samples routinely distributed for testing to determine accuracy? | Yes
229 | 8.13.2 Do PT samples come from providers who are accredited or approved? | No | CD4 not approved
230 | 8.13.3 Are PT specimens handled and tested the same way as patient specimens? | Yes
231 | 8.13.4 Is cause analysis performed for unacceptable PT results? | No | No corrective action for failed EQA
232 | 8.13.5 Is corrective action documented for unacceptable PT results? | No | No documentation
233 | 8.14 Are test requests checked with test results, thereby assuring the accuracy and completion of all tests? | Partial | No clinical information
234 | 9.1 Test Result Reporting System | Yes
235 | 9.2 Testing Personnel | Partial | Some signatures no traceable
236 | 9.3 Test Result Records | Yes
237 | 9.4 Analytic System/Method Tracing | Yes
238 | 9.5 Result Cross-check System | Partial | Transcribed results are not always reviewed by head of section
239 | 9.6 Archived Data Labeling and Storage | Partial | Archived data is stored in non-locked cabinets in hematology and serology
240 | 9.7 Information and Data Backup System | Partial | Procedure is not consistent with correct practice
241 | 9.8 Test Result ReportIs the laboratory result report(s) in a standard form determined to be acceptable by its customers?
242 | 9.8.1 Is the laboratory issuing the report clearly identified? | Yes
243 | 9.8.2 Does the report contain the patient's name, address, and the hospital/destination of the report? | Yes
244 | 9.8.3 Is the name of the person requesting the test indicated on the report? | Yes
245 | 9.8.4 Is the type of sample received and the test requested included in the report? | Yes
246 | 9.8.5 Are the date and time for specimen collection, receipt of specimen, and release of report indicated? | Yes
247 | 9.8.6 Does the report indicate biological reference ranges for each test? | Yes
248 | 9.8.7 Is the result reported in SI units where applicable? | Yes
249 | 9.8.8 Is there space for interpretation of results, when applicable, and for indication of when specimens are received and unsuitable for the procedure requested for testing? | Yes
250 | 9.8.9 Does the result contain the name of the person authorizing release of the report and the signature of the person accepting responsibility for its content? | Yes
251 | 9.9 Test Result | Yes
252 | 10.1 Are all laboratory-documented occurrence reports indicating the root cause of the problem(s) and corrective & preventive actions taken to prevent recurrence? (There must be at least a description of what happened and what was done to prevent it from happening again.) | Partial | No corrective action for unacceptable EQA results
253 | 10.2 Is non-conforming work reviewed and submitted for troubleshooting and cause analysis? | Partial | No evidence
254 | 10.3 Is corrective action performed on all non-conforming aspects of the quality management system documented?
255 | 10.3.1 Are results withheld, if indicated by the level of control violated? (ISO 4.9.1 part d) | No | Results not withheld
256 | 10.3.2 Have these been recalled and corrected, if results have been released? (ISO 4.9.1 part f) | Yes
257 | 10.3.3 Is this approved by an authorized person, when testing resumes? (ISO 4.9.1.part g) | No | Without performance IQC
258 | 10.4 Are discordant results tracked and appropriate corrective action taken? | Partial | Discordant AFB result between CPGH and TB reference lab not resolved yet, Sample ID - 105011
259 | 11.1 Are graphical tools (charts and graphs) used to communicate quality findings and identify trends? | Partial | Not used to communicate quality indicators
260 | 11.2 Are quality indicators (TAT, rejected specimens, stock outs, etc.) selected, tracked, and reviewed regularly to monitor laboratory performance and identify potential quality improvement activities? | Yes | TAT, Specimen rejection and customer satisfaction
261 | 11.3 Are the outcomes of internal and external audits, PT, customer feedback and all other information derived from the tracking of quality indicators used to improve lab performance? | Partial | Outcome not used for quality improvement
262 | 11.4 Is the outcome of the action taken checked and monitored to determine the effectiveness of improved quality of lab performance? | No
263 | 12.1 Is the size of the laboratory adequate and the layout of the laboratory, as a whole, organized so that workstations are positioned for optimal workflow? | Yes
264 | 12.2 Are the patient care and testing areas of the laboratory distinctly separate from one another? | Partial | Casualty lab not directly separated
265 | 12.3 Is each individual workstation maintained free of clutter and set up for efficient operation?
266 | 12.3.1 Does the equipment placement/layout facilitate optimum workflow? | Yes
267 | 12.3.2 Are all needed supplies present and easily accessible? | Yes
268 | 12.3.3 Are the chairs/stools at the workstations appropriate for bench height and the testing operations being performed? (ISO 15190: 6.3.5) | No | Torn chairs in microbiology labs
269 | 12.3.4 Is reference material readily available ( critical values and required action, population reference ranges, frequently called numbers? | No | Critical values have not been identified - Microbiology section
270 | 12.4 Is the physical work environment appropriate for testing?
271 | 12.4.1 Free of clutter | Yes
272 | 12.4.2 Adequately ventilated? | Yes
273 | 12.4.3 Free of excess moisture? | Yes
274 | 12.4.4 Adequately lit? | Yes
275 | 12.4.5 Climate-controlled for optimum equipment function? | Yes
276 | 12.4.6 Are filters checked, cleaned and/or replaced at regular intervals, where air-conditioning is installed? | Yes
277 | 12.4.7 Are wires and cables properly located and protected from traffic? | Yes
278 | 12.4.8 Is there a functioning back-up power supply (generator)? | Yes
279 | 12.4.9 Is critical equipment supported by uninterrupted power source (UPS) systems? | Yes
280 | 12.4.10 Is equipment placed appropriately (away from water hazards, out of traffic areas)? | Yes
281 | 12.4.11 Is a contingency plan in place for continued testing in the event of prolonged electricity disruption? | Yes
282 | 12.4.12 Are appropriate provisions made for adequate water supply, including deionized water (DI) or distilled water, if needed? | Yes
283 | 12.4.13 Is clerical work completed outside the testing area? | Yes
284 | 12.4.14 Is major safety signage posted and enforced including NO EATING, SMOKING, DRINKING? | Yes
285 | 12.5 Is the laboratory properly secured from unauthorized access with appropriate signage? | Yes
286 | 12.6 Is laboratory-dedicated cold and room temperature storage free of staff food items, and are patient samples stored separately from reagents and blood products in the laboratory refrigerators and freezers? | Yes
287 | 12.7 Is the work area clean and free of leakage & spills, and are disinfection procedures conducted and documented? | Yes
288 | 12.8 Is a certified and appropriate Biosafety cabinet (or an acceptable alternative processing procedure) in use for all specimens or organisms considered to be highly contagious by airborne routes? (Biosafety cabinet should be recertified according to national protocol). | Yes
289 | 12.9 Is a laboratory safety manual available, accessible, and up-to-date?
290 | 12.9.1 Blood and Body Fluid Precautions | Yes
291 | 12.9.2 Hazardous Waste Disposal | Yes
292 | 12.9.3 Hazardous Chemicals / Materials | Yes
293 | 12.9.4 MSDS Sheets | Yes
294 | 12.9.5 Personal protective equipment | Yes
295 | 12.9.6 Vaccination | Yes
296 | 12.9.7 Post-Exposure Prophylaxis | Yes
297 | 12.9.8 Fire Safety | Yes
298 | 12.9.9 Electrical safety | Yes
299 | 12.10 Is sufficient waste disposal available and is waste separated into infectious and non-infectious waste, with infectious waste autoclaved, incinerated, or buried? | Partial | Used culture plates are not auto-cleaned before disposal
300 | 12.11 Are hazardous chemicals / materials properly handled?
301 | 12.11.1 Are hazardous chemicals properly labeled? | Yes
302 | 12.11.2 Are hazardous chemicals properly stored? | Yes
303 | 12.11.3 Are hazardous chemicals properly utilized? | Yes
304 | 12.11.4 Are hazardous chemicals properly disposed? | No | Xylene disposed off in the sink
305 | 12.12 Are 'sharps' handled and disposed of properly in 'sharps' containers that are appropriately utilized? | Yes
306 | 12.13 Is fire safety included as part of the laboratory’s overall safety program?
307 | 12.13.1 Are all electrical cords, plugs, and receptacles used appropriately and in good repair? | Yes
308 | 12.13.2 Is an appropriate fire extinguisher available, properly placed, in working condition, and routinely inspected? | No | Extinguisher last serviced - 2010, no checks by staff
309 | 12.13.3 Is an operational fire warning system in place in laboratory with periodic fire drills? | No | No fire drill done
310 | 12.14 Are safety inspections or audits conducted regularly and documented? | Yes
311 | 12.15 Is standard safety equipment available and in use in the laboratory?
312 | 12.15.1 Biosafety cabinet(s) | Yes
313 | 12.15.2 Covers on centrifuge(s) | Yes
314 | 12.15.3 Hand-washing station | Yes
315 | 12.15.4 Eyewash station/bottle(s) and showers where applicable | Yes
316 | 12.15.5 Spill kit(s) | Yes
317 | 12.15.6 First aid kit(s) | Yes
318 | 12.16 Is personal protective equipment (PPE) easily accessible at the workstation and utilized appropriately and consistently? | Yes
319 | 12.17 Are laboratory personnel offered appropriate vaccination/preventive measures? | Partial | No vaccination record for Abigail Slegi PF13
320 | 12.18 Are post-exposure prophylaxis policies and procedures posted and implemented after possible and known exposures? | Yes
321 | 12.19 Are occupational injuries, medical screening or illnesses documented in the safety occurrence log? | Yes
322 | 12.10 Are drivers/couriers and cleaners working with the laboratory trained in Biosafety practices relevant to their job tasks? | Partial | Cleaners have not been trained
323 | 12.21 Is a trained safety officer designated to implement and monitor the safety program in the laboratory, including the training of other staff? | Yes

## Scores
Question | Description | Points
1 | 1.1 | 4
8 | 1.2 | 1
9 | 1.3 | 1
10 | 1.4 | 1
43 | 1.5 | 2
44 | 1.6 | 1
45 | 1.7 | 1
46 | 1.8 | 2
47 | 1.9 | 2
48 | 1.1 | 0
49 | 2.1 | 2
50 | 2.2 | 1
74 | 2.3 | 1
98 | 2.4 | 3
99 | 2.5 | 2
100 | 3.1 | 2
101 | 3.2 | 2
102 | 3.3 | 2
103 | 3.4 | 3
104 | 3.5 | 3
118 | 3.6 | 1
119 | 3.7 | 1
120 | 3.8 | 1
132 | 4.1 | 2
133 | 4.2 | 2
134 | 4.3 | 2
135 | 4.4 | 1
136 | 5.1 | 2
137 | 5.2 | 1
138 | 5.3 | 1
146 | 5.4 | 1
154 | 5.5 | 2
155 | 5.6 | 1
156 | 5.7 | 1
157 | 5.8 | 1
158 | 5.9 | 2
159 | 5.1 | 1
160 | 5.11 | 1
161 | 5.12 | 1
162 | 5.13 | 2
163 | 5.14 | 2
164 | 5.15 | 1
165 | 6.1 | 1
170 | 6.2 | 1
171 | 7.1 | 2
172 | 7.2 | 2
173 | 7.3 | 2
174 | 7.4 | 2
175 | 7.5 | 2
176 | 7.6 | 1
177 | 7.7 | 1
181 | 7.8 | 2
182 | 7.9 | 2
183 | 7.1 | 2
184 | 7.11 | 2
195 | 7.12 | 2
196 | 7.13 | 1
197 | 7.14 | 0
198 | 7.15 | 1
199 | 8.1 | 2
200 | 8.2 | 1
211 | 8.3 | 2
213 | 8.4 | 2
214 | 8.5 | 2
215 | 8.6 | 3
216 | 8.7 | 2
217 | 8.8 | 1
218 | 8.9 | 1
219 | 8.1 | 3
220 | 8.11 | 2
226 | 8.12 | 2
227 | 8.13 | 1
233 | 8.14 | 1
234 | 9.1 | 2
235 | 9.2 | 1
236 | 9.3 | 2
237 | 9.4 | 2
238 | 9.5 | 1
239 | 9.6 | 1
240 | 9.7 | 1
241 | 9.8 | 2
251 | 9.9 | 2
252 | 10.1 | 1
253 | 10.2 | 1
254 | 10.3 | 1
258 | 10.4 | 1
259 | 11.1 | 1
260 | 11.2 | 5
261 | 11.3 | 1
262 | 11.4 | 0
263 | 12.1 | 2
264 | 12.2 | 1
265 | 12.3 | 1
270 | 12.4 | 2
285 | 12.5 | 2
286 | 12.6 | 2
287 | 12.7 | 2
288 | 12.8 | 2
289 | 12.9 | 3
299 | 12.1 | 1
300 | 12.1 | 1
305 | 12.1 | 2
306 | 12.1 | 1
310 | 12.1 | 2
311 | 12.1 | 2
318 | 12.1 | 2
319 | 12.1 | 1
320 | 12.1 | 2
321 | 12.1 | 2
322 | 12.1 | 1
323 | 12.2 | 2

## Summary of Assessment Findings
Field | Value
Commendations
Challenges
Recommendations

## Action Plan
Action | Incharge | Timeline